## Sheet1
2023-10-25 | JB
Raysafe X2
LA5
Raysafe at iso, G=90, extend kVs, take series of measurements to then find HVL and dose in-air profile for each combo
Service mode; XI tab, Acquisition, kV, Dynamic gain
PU--> kVSColl--> extend filter & fan (Foil & shape |  |  |  |  | Foil | 0 - out | 1 - in
[5 empty] | Shape | 0 none | 1 - half fan | 2 - full fan
SPOTLIGHT | Blades(x,y) (14,14,10.7,10.7)
125kV | Full Fan | Ti Filter in | [9 empty] | Pelvis
40mA | 40mS | Focal Spot Large |  |  |  |  | 120kV - no bow tie, CAX (40mA,40ms, large focal spot): |  |  |  |  | 125kV | Blades(24.7,3.3,10.7,10.7) |  |  | Half Fan
[7 empty] | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl) | CAX:
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl) |  | 120.5 | 6.523 |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.118 | 3 | [8 empty] | 125.7 | 106.3 | 2.633 | 8.4
-14 |  | 0.134 | 3.445
-12 |  | 0.178 | 4.45 | [8 empty] | 125 Full Fan CAX (Spotlight, blades (14,14,10.7,10.7)
-10 |  | 0.522 | 13.12 | [8 empty] | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
-8 |  | 1.73 | 43.6 | [8 empty] | 124.5 | 107.1 | 2.645 | 8.42
-6 |  | 3.178 | 79.82 | 3.84
-4 |  | 4.386 | 109.8 | 3.72
-2 |  | 4.576 | 114.2 | 3.74
0 |  | 3.737 | 93.88 | 3.87
2 |  | 2.291 | 57.54 | 4.05
4 |  | 0.986 | 24.72
6 |  | 0.225 | 5.641
8 |  | 0.151 | 3.73
10 |  | 0.123 | 3.093
12 |  | 0.049 | 1.136
14 |  | collision detection
16
HEAD | Blades (14,14,10.7,10.7) | [5 empty] | 100kV - no bow tie, CAX (40mA,40ms, large focal spot):
100kV | Full Fan | Ti Filter in |  |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
40mA | 40mS | Focal Spot Large |  |  |  |  | 100.4 | 112.8 | 3.022 | 5.04 | 4.7
distance in crossplane
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.824 | 20.57
-14 |  | 3.993 | 99.68 | 9.98
-12 |  | 7.645 | 191.1 | 9.81
-10 |  | 8.412 | 209.6 | 9.65
-8 |  | 12.45 | 311.3 | 9.4
-6 |  | 30.7 | 760.3 | 8.56
-4 |  | 44.9 | 1.112 | 7.88
-2 |  | 54.68 | 1.352 | 7.49 | 14
0 | 100.6 | 59.01 | 1.459 | 7.37 | 13
2 |  | 55.33 | 1.37 | 7.45 | 13
4 |  | 45.22 | 1.12 | 7.81
6 |  | 31.3 | 776.4 | 8.48
8 |  | 13.16 | 329.5 | 9.42
10
12
14 |  | collision detection
16
IMAGE GENTLY | Blades (14,14,10.7,10.7) | [5 empty] | 80kV - no bow tie, CAX (40mA,40ms, large focal spot):
80kV | Full Fan | Ti Filter in |  |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
40mA | 40mS | Focal Spot Large | [5 empty] | 28.89 | 717 | 6.18
distance in crossplane
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.35 | 8.9
-14 |  | 1.384 | 34.82 | 8.09
-12 |  | 2.695 | 67.69 | 8.12
-10 |  | 3.013 | 75.57 | 8.05
-8 |  | 4.7 | 117.9 | 7.95
-6 |  | 13.21 | 331.8 | 7.11
-4 |  | 20.93 | 521.6 | 6.62
-2 |  | 26.76 | 666 | 6.26
0 |  | 29.92 | 720.9 | 6.17
2 |  | 26.91 | 667.5 | 6.2
4 |  | 21.1 | 526 | 6.61
6 |  | 13.5 | 339.1 | 7.14
8 |  | 4.948 | 123.5 | 8.07
10 |  | collision detection
12
14 |  | collision detection
16

## Collated results
All DynamicGain kV shots
| 2023-10-25 |  |  | 2023-10-25 | [8 empty] | 2023-10-25 |  |  | 2022-08-01 |  |  |  |  | 2022-08-01
Protocol | Head |  |  | Spotlight | [8 empty] | Img Gently |  |  | Pelvis |  |  |  |  | Pelvis Large
kV | 100 kV |  |  | 125 kV | [8 empty] | 80 kV |  |  | 125 kV |  |  |  |  | 140 kV
Fan | Full |  |  | Full | [8 empty] | Full |  |  | Half |  |  |  |  | Half
Ti filter | Yes |  |  | Yes | [8 empty] | Yes |  |  | Yes |  |  |  |  | Yes
Blades | (14,14,10.7,10.7) |  |  | (14,14,10.7,10.7) | [8 empty] | (14,14,10.7,10.7) |  |  | (24.7,3.3,10.7,10.7) |  |  |  |  | (24.7,3.3,10.7,10.7) |  |  | 1350
mA/rdg | 40 | *not clinical setting |  | 40 | *not clinical setting | [7 empty] | 40 | *not clinical setting |  | 60 |  |  |  |  | 75
mS/rdg | 40 |  |  | 40 | [8 empty] | 40 |  |  | 20 |  |  |  |  | 25
Focal spot | Large |  |  | Large | [8 empty] | Large |  |  | Large |  |  |  |  | Large
|  |  |  | frames/s | 15
|  |  |  | s/frame | 0.066667 | s
Crossplane | [16 empty] | interpolated from meas |  |  |  |  | interpolated from meas
/cm | Dose uGy | uGy/s | HVL (mm Al) | Dose uGy | uGy/s | HVL (mm Al) | uGy/s Coast | HVL (mm Al) Coast | Scaling factor | Air kerma per projection(mGy/frame) |  |  | Dose uGy | uGy/s | HVL (mm Al) | Dose uGy | uGy/s | HVL (mm Al) | uGy/s Coast | HVL (mm Al) coast | Dose uGy | uGy/s | HVL (mm Al)
-16 | 0.824 | 20.57 |  | 0.118 |  |  | using continuous mode | [5 empty] | 0.35 | 8.9
-14 | 3.993 | 99.68 | 9.98 | 0.134 | [6 empty] | uGy/s |  | 1.384 | 34.82 | 8.09 |  | 42.3 | 8.589 |  |  |  | 72.78 | 9.751
-12 | 7.645 | 191.1 | 9.81 | 0.178 | 3.2225 |  |  |  |  | 0.0026068 | 39.1016 |  | 2.695 | 67.69 | 8.12 |  | 195.8 | 10.571 |  |  |  | 352.9 | 11.333
-10 | 8.412 | 209.6 | 9.65 | 0.522 | 3.9475 |  |  |  |  | 0.0031932 | 47.8987 |  | 3.013 | 75.57 | 8.05 |  | 340.6 | 9.958 |  |  |  | 605.5 | 10.759
-8 | 12.45 | 311.3 | 9.4 | 1.73 | 8.785 |  |  |  |  | 0.0071064 | 106.5966 |  | 4.7 | 117.9 | 7.95 |  | 557.6 | 9.204 |  |  |  | 963.7 | 9.893
-6 | 30.7 | 760.3 | 8.56 | 3.178 | 28.36 |  |  |  |  | 0.022941 | 344.1182 | 9.1179 | 13.21 | 331.8 | 7.11 |  | 739.7 | 8.809 |  |  |  | 1249 | 9.431
-4 | 44.9 | 1112 | 7.88 | 4.386 | 61.71 | 3.84 |  |  |  | 0.049919 | 748.7847 | 8.6451 | 20.93 | 521.6 | 6.62 |  | 878.6 | 8.353 |  |  |  | 1471 | 8.993
-2 | 54.68 | 1352 | 7.49 | 4.576 | 94.81 | 3.72 |  |  |  | 0.076695 | 1150.4178 | 8.375 | 26.76 | 666 | 6.26 |  | 971.5 | 8.146 |  |  |  | 1639 | 8.645
0 | 59.01 | 1459 | 7.37 | 3.737 | 112 | 3.74 | 1359 | 8.42 | 12.1339 | 0.0906 | 1359 | 8.375 | 29.92 | 720.9 | 6.17 |  | 1018 | 8.06 | 1080 | 8.39 |  | 1706 | 8.539
2 | 55.33 | 1370 | 7.45 | 2.291 | 104.04 | 3.87 |  |  |  | 0.084161 | 1262.4139 | 8.42 | 26.91 | 667.5 | 6.2 |  | 1024 | 8.044 |  |  |  | 1727 | 8.627
4 | 45.22 | 1120 | 7.81 | 0.986 | 75.71 | 4.05 |  |  |  | 0.061244 | 918.6597 | 8.7127 | 21.1 | 526 | 6.61 |  | 1039 | 8.103 |  |  |  | 1730 | 8.582
6 | 31.3 | 776.4 | 8.48 | 0.225 | 41.13 |  |  |  |  | 0.033271 | 499.0685 | 9.1179 | 13.5 | 339.1 | 7.14 |  | 1029 | 7.905 |  |  |  | 1720 | 8.638
8 | 13.16 | 329.5 | 9.42 | 0.151 | 15.1805 |  |  |  |  | 0.01228 | 184.1991 |  | 4.948 | 123.5 | 8.07 |  | 1007 | 8.17 |  |  |  | 1704 | 8.592
10 | 8.412 | 209.6 | 9.65 | 0.123 | 4.6855 |  |  |  |  | 0.0037902 | 56.8535 |  |  | 123.5 | 8.07 |  | 1008 | 7.891 |  |  |  | 1695 | 8.551
12 | 7.645 | 191.1 | 9.81 | 0.049 | 3.4115 |  |  |  |  | 0.0027597 | 41.3949 |  |  | 67.69 | 8.05 |  | 1008 | 7.891 |  |  |  | 1695 | 8.551
14 | 3.993 | 99.68 | 9.98 | 0.134 | [9 empty] | 34.82 | 8.12 |  | 1008 | 7.891 |  |  |  | 1695 | 8.551
16 | 0.824 | 20.57 |  | 0.118 | [9 empty] | 8.9 | 8.09 |  | 1008 | 7.891 |  |  |  | 1695 | 8.551
|  | *extrapolated |  | *extrapolated | [9 empty] | *extrapolated |  |  | *extrapolated |  |  |  |  | *extrapolated
Note on J17: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Interpolated from previous sheet for a 1cm asymmetry - setup error