## Sheet1
2023-10-25 | JB
Raysafe X2
LA5
Raysafe at iso, G=90, extend kVs, take series of measurements to then find HVL and dose in-air profile for each combo
Service mode; XI tab, Acquisition, kV, Dynamic gain
PU--> kVSColl--> extend filter & fan (Foil & shape |  |  |  |  | Foil | 0 - out | 1 - in
[5 empty] | Shape | 0 none | 1 - half fan | 2 - full fan
SPOTLIGHT | Blades(x,y) (14,14,10.7,10.7)
125kV | Full Fan | Ti Filter in | [9 empty] | Pelvis
40mA | 40mS | Focal Spot Large |  |  |  |  | 120kV - no bow tie, CAX (40mA,40ms, large focal spot): |  |  |  |  | 125kV | Blades(24.7,3.3,10.7,10.7) |  |  | Half Fan
[7 empty] | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl) | CAX:
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl) |  | 120.5 | 6.523 |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.118 | 3 | [8 empty] | 125.7 | 106.3 | 2.633 | 8.4
-14 |  | 0.134 | 3.445
-12 |  | 0.178 | 4.45 | [8 empty] | 125 Full Fan CAX (Spotlight, blades (14,14,10.7,10.7)
-10 |  | 0.522 | 13.12 | [8 empty] | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
-8 |  | 1.73 | 43.6 | [8 empty] | 124.5 | 107.1 | 2.645 | 8.42
-6 |  | 3.178 | 79.82 | 3.84
-4 |  | 4.386 | 109.8 | 3.72
-2 |  | 4.576 | 114.2 | 3.74
0 |  | 3.737 | 93.88 | 3.87
2 |  | 2.291 | 57.54 | 4.05
4 |  | 0.986 | 24.72
6 |  | 0.225 | 5.641
8 |  | 0.151 | 3.73
10 |  | 0.123 | 3.093
12 |  | 0.049 | 1.136
14 |  | collision detection
16
HEAD | Blades (14,14,10.7,10.7) | [5 empty] | 100kV - no bow tie, CAX (40mA,40ms, large focal spot):
100kV | Full Fan | Ti Filter in |  |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
40mA | 40mS | Focal Spot Large |  |  |  |  | 100.4 | 112.8 | 3.022 | 5.04 | 4.7
distance in crossplane
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.824 | 20.57
-14 |  | 3.993 | 99.68 | 9.98
-12 |  | 7.645 | 191.1 | 9.81
-10 |  | 8.412 | 209.6 | 9.65
-8 |  | 12.45 | 311.3 | 9.4
-6 |  | 30.7 | 760.3 | 8.56
-4 |  | 44.9 | 1.112 | 7.88
-2 |  | 54.68 | 1.352 | 7.49 | 14
0 | 100.6 | 59.01 | 1.459 | 7.37 | 13
2 |  | 55.33 | 1.37 | 7.45 | 13
4 |  | 45.22 | 1.12 | 7.81
6 |  | 31.3 | 776.4 | 8.48
8 |  | 13.16 | 329.5 | 9.42
10
12
14 |  | collision detection
16
IMAGE GENTLY | Blades (14,14,10.7,10.7) | [5 empty] | 80kV - no bow tie, CAX (40mA,40ms, large focal spot):
80kV | Full Fan | Ti Filter in |  |  |  |  | kVp | Dose | mGy/s | HVL (mm Al) | TF (mmAl)
40mA | 40mS | Focal Spot Large | [5 empty] | 28.89 | 717 | 6.18
distance in crossplane
| kVp | Dose uGy | uGy/s | HVL (mm Al) | TF (mmAl)
-16 |  | 0.35 | 8.9
-14 |  | 1.384 | 34.82 | 8.09
-12 |  | 2.695 | 67.69 | 8.12
-10 |  | 3.013 | 75.57 | 8.05
-8 |  | 4.7 | 117.9 | 7.95
-6 |  | 13.21 | 331.8 | 7.11
-4 |  | 20.93 | 521.6 | 6.62
-2 |  | 26.76 | 666 | 6.26
0 |  | 29.92 | 720.9 | 6.17
2 |  | 26.91 | 667.5 | 6.2
4 |  | 21.1 | 526 | 6.61
6 |  | 13.5 | 339.1 | 7.14
8 |  | 4.948 | 123.5 | 8.07
10 |  | collision detection
12
14 |  | collision detection
16

## Collated results
All DynamicGain kV shots
| 2023-10-25 | [6 empty] | 2023-10-25 | [5 empty] | 2023-10-25 | [5 empty] | 2022-08-01 |  |  | 2022-08-01
Protocol | Head | [6 empty] | Spotlight | [5 empty] | Img Gently | [5 empty] | Pelvis |  |  | Pelvis Large
kV | 100 kV | [6 empty] | 125 kV | [5 empty] | 80 kV | [5 empty] | 125 kV |  |  | 140 kV
Fan | Full | [6 empty] | Full | [5 empty] | Full | [5 empty] | Half |  |  | Half
Ti filter | Yes | [6 empty] | Yes | [5 empty] | Yes | [5 empty] | Yes |  |  | Yes
Blades | (14,14,10.7,10.7) |  | Clinical Setting: |  |  |  |  | (14,14,10.7,10.7) |  | Clinical Setting: |  |  |  | (14,14,10.7,10.7) |  | Clinical Setting: |  |  |  | (24.7,3.3,10.7,10.7) |  |  | (24.7,3.3,10.7,10.7)
mA/rdg | 40 |  | mA | 15 |  |  |  | 40 |  | mA | 60 |  |  | 40 |  | mA | 10 |  |  | 60 |  |  | 75
mS/rdg | 40 |  | ms | 20 |  |  |  | 40 |  | ms | 25 |  |  | 40 |  | ms | 20 |  |  | 20 |  |  | 25
Focal spot | Large |  | frames/s | 15 |  |  |  | Large |  | frames/s | 15 |  |  | Large |  | frames/s | 15 |  |  | Large |  |  | Large
[5 empty] | uGy/s if contstant output | [15 empty] | FROM PREV MEASUREMENTS |  |  | FROM PREV MEASUREMENTS
Crossplane | [7 empty] | FROM DEVICE | [12 empty] | interpolated from meas |  |  | interpolated from meas
/cm | Dose uGy | uGy/s | HVL (mm Al) |  |  | uGy/mAs | uGy/s for clin set | Dose uGy | uGy/s | HVL (mm Al) |  | uGy/mAs | uGy/s for clin set | Dose uGy | uGy/s | HVL (mm Al) |  | uGy/mAs | uGy/s for clin set | Dose uGy | uGy/s | HVL (mm Al) | Dose uGy | uGy/s | HVL (mm Al)
-16 | 0.824 | 20.57 |  |  | 20.6 | 0.515 | 2.3175 | 0.118 |  |  |  | 0.07375 | 1.6594 | 0.35 | 8.9 |  |  | 0.21875 | 0.65625
-14 | 3.993 | 99.68 | 9.98 |  | 99.825 | 2.4956 | 11.2303 | 0.134 |  |  |  | 0.08375 | 1.8844 | 1.384 | 34.82 | 8.09 |  | 0.865 | 2.595 |  | 42.3 | 8.589 |  | 72.78 | 9.751
-12 | 7.645 | 191.1 | 9.81 |  | 191.125 | 4.7781 | 21.5016 | 0.178 | 3.2225 |  |  | 0.11125 | 2.5031 | 2.695 | 67.69 | 8.12 |  | 1.6844 | 5.0531 |  | 195.8 | 10.571 |  | 352.9 | 11.333
-10 | 8.412 | 209.6 | 9.65 |  | 210.3 | 5.2575 | 23.6588 | 0.522 | 3.9475 |  |  | 0.32625 | 7.3406 | 3.013 | 75.57 | 8.05 |  | 1.8831 | 5.6494 |  | 340.6 | 9.958 |  | 605.5 | 10.759
-8 | 12.45 | 311.3 | 9.4 |  | 311.25 | 7.7812 | 35.0156 | 1.73 | 8.785 |  |  | 1.0812 | 24.3281 | 4.7 | 117.9 | 7.95 |  | 2.9375 | 8.8125 |  | 557.6 | 9.204 |  | 963.7 | 9.893
-6 | 30.7 | 760.3 | 8.56 |  | 767.5 | 19.1875 | 86.3438 | 3.178 | 28.36 |  |  | 1.9862 | 44.6906 | 13.21 | 331.8 | 7.11 |  | 8.2562 | 24.7688 |  | 739.7 | 8.809 |  | 1249 | 9.431
-4 | 44.9 | 1112 | 7.88 |  | 1122.5 | 28.0625 | 126.2812 | 4.386 | 61.71 | 3.84 |  | 2.7412 | 61.6781 | 20.93 | 521.6 | 6.62 |  | 13.0812 | 39.2437 |  | 878.6 | 8.353 |  | 1471 | 8.993
-2 | 54.68 | 1352 | 7.49 |  | 1367 | 34.175 | 153.7875 | 4.576 | 94.81 | 3.72 |  | 2.86 | 64.35 | 26.76 | 666 | 6.26 |  | 16.725 | 50.175 |  | 971.5 | 8.146 |  | 1639 | 8.645
0 | 59.01 | 1459 | 7.37 |  | 1475.25 | 36.8812 | 165.9656 | 3.737 | 112 | 3.74 |  | 2.3356 | 52.5516 | 29.92 | 720.9 | 6.17 |  | 18.7 | 56.1 |  | 1018 | 8.06 |  | 1706 | 8.539
2 | 55.33 | 1370 | 7.45 |  | 1383.25 | 34.5812 | 155.6156 | 2.291 | 104.04 | 3.87 |  | 1.4319 | 32.2172 | 26.91 | 667.5 | 6.2 |  | 16.8187 | 50.4562 |  | 1024 | 8.044 |  | 1727 | 8.627
4 | 45.22 | 1120 | 7.81 |  | 1130.5 | 28.2625 | 127.1812 | 0.986 | 75.71 | 4.05 |  | 0.61625 | 13.8656 | 21.1 | 526 | 6.61 |  | 13.1875 | 39.5625 |  | 1039 | 8.103 |  | 1730 | 8.582
6 | 31.3 | 776.4 | 8.48 |  | 782.5 | 19.5625 | 88.0312 | 0.225 | 41.13 |  |  | 0.14062 | 3.1641 | 13.5 | 339.1 | 7.14 |  | 8.4375 | 25.3125 |  | 1029 | 7.905 |  | 1720 | 8.638
8 | 13.16 | 329.5 | 9.42 |  | 329 | 8.225 | 37.0125 | 0.151 | 15.1805 |  |  | 0.094375 | 2.1234 | 4.948 | 123.5 | 8.07 |  | 3.0925 | 9.2775 |  | 1007 | 8.17 |  | 1704 | 8.592
10 | 8.412 | 209.6 | 9.65 |  | 210.3 | 5.2575 | 23.6588 | 0.123 | 4.6855 |  |  | 0.076875 | 1.7297 |  | 123.5 | 8.07 |  | 0 | 0 |  | 1008 | 7.891 |  | 1695 | 8.551
12 | 7.645 | 191.1 | 9.81 |  | 191.125 | 4.7781 | 21.5016 | 0.049 | 3.4115 |  |  | 0.030625 | 0.68906 |  | 67.69 | 8.05 |  | 0 | 0 |  | 1008 | 7.891 |  | 1695 | 8.551
14 | 3.993 | 99.68 | 9.98 |  | 99.825 | 2.4956 | 11.2303 | 0.134 |  |  |  | 0.08375 | 1.8844 |  | 34.82 | 8.12 |  | 0 | 0 |  | 1008 | 7.891 |  | 1695 | 8.551
16 | 0.824 | 20.57 |  |  | 20.6 | 0.515 | 2.3175 | 0.118 |  |  |  | 0.07375 | 1.6594 |  | 8.9 | 8.09 |  | 0 | 0 |  | 1008 | 7.891 |  | 1695 | 8.551
|  | *extrapolated | [5 empty] | *extrapolated | [6 empty] | *extrapolated | [5 empty] | *extrapolated |  |  | *extrapolated
Note on N16: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Interpolated from previous sheet for a 1cm asymmetry - setup error