## Initial tests
[5 empty] | Head mode | [6 empty] | 100 | kv | hVL | 8.5 mmAl
[5 empty] | test 2 |  |  |  | Blade pos | x1 | 14 | 15 | mA
couch (cm) | kv | Gy | uGy/s |  | couch (cm) | uGy/s |  |  |  | x2 | 14 | 20 | ms
0 |  |  | 201 |  | 0 |  |  |  |  | y1 | 10.7
-5 |  |  | 199 |  | 0.5 |  |  |  |  | y2 | 10.7
-10 |  |  | 59.5 |  | 1
-15 |  |  | 0 |  | 1.5
2
2.5
3
3.5
4
4.5
5
5.5
6
6.5
7
7.5
8
8.5
9
9.5
10
10.5
11
11.5
12
12.5
13
13.5
14
14.5
15
|  |  |  | 0 | 0 | 167 |  | 165
|  |  |  | -0.5 | 999.5 | 165
|  |  |  | -1 | 999 | 164
|  |  |  | -1.5 | 998.5 | 162
|  |  |  | -2 | 998 | 156
|  |  |  | -2.5 | 997.5 | 151
|  |  |  | -3 | 997 | 145
|  |  |  | -3.5 | 996.5 | 136
|  |  |  | -4 | 996 | 133
|  |  |  | -4.5 | 995.5 | 111
|  |  |  | -5 | 995 | 105
|  |  |  | -5.5 | 994.5 | 94
|  |  |  | -6 | 994 | 83 |  | 72.8
|  |  |  | -6.5 | 993.5 | 72 |  | 60.4
|  |  |  | -7 | 993 | 59 |  | 46.6
|  |  |  | -7.5 | 992.5 | 44 |  | 33
|  |  |  | -8 | 992 |  | rf low? | 28.5
|  |  |  | -8.5 | 991.5 |  |  | 26
|  |  |  | -9 | 991 |  |  | 24.5
|  |  |  | -9.5 | 990.5 |  |  | 24.5
|  |  |  | -10 | 990 |  |  | 23.5
|  |  |  | -10.5 | 989.5 |  |  | 22
|  |  |  | -11 | 989 |  |  | 21.7
|  |  |  | -11.5 | 988.5 |  |  | 21.4
|  |  |  | -12 | 988 |  |  | 21
|  |  |  | -12.5 | 987.5 |  |  | 20.7
|  |  |  | -13 | 987 |  |  | 17.9
|  |  |  | -13.5 | 986.5 |  |  | 17.9
|  |  |  | -14 | 986
|  |  |  | -14.5 | 985.5
|  |  |  | -15 | 985

## Head
Head |  |  | x1 | 14
| kvp | 100 | x2 | 14
| mA | 15 | y1 | 10.7
| ms | 20 | y2 | 10.7
| frames/s | 15
| Fan type | Full
| Arc type | Half
kV | 100.38 | kV
HVL | 6.887 | mm Al
Total filt | 10.829 | mmAl TF
Dose | 1.392 | mGy
Dose rate | 0.171 | mGy/s |  | 0.011594
frame/s | 14.7489 | frame/s

## Pelvis
Pelvis |  |  | x1 | 24.7 | cm | rf high
| kvp | 125 | x2 | 3.3
| mA | 60 | y1 | 10.7
| ms | 20 | y2 | 10.7
| frames/s | 15
| Fan type | Half
| Arc type | Full
kV | 129.04 | kV
HVL | 8.371 | mm Al
Total filt | - | mmAl TF
Dose | 9.28 | mGy
Dose rate | 1.11 | mGy/s
frame/s | 14.7574 | frame/s

## Pelvis large
Pelvis large |  |  | x1 | 24.7 | cm | rf high
| kvp | 140 | x2 | 3.3
| mA | 75 | y1 | 10.7
| ms | 25 | y2 | 10.7
| frames/s | 5
| Fan type | Half
| Arc type | Full
kV | 143.29 | kV
HVL | 8.942 | mm Al
Total filt | - | mmAl TF
Dose | 6.008 | mGy
Dose rate | 0.75 | mGy/s
frame/s | 4.9669 | frame/s

## Breast
Breast 360 | half arc | full fan | Breast 200 | half arc | full fan | x1 | 14
| kvp | 125 |  | kvp | 125 | x2 | 14
| mA | 10 |  | mA | 10 | y1 | 10.7
| ms | 10 |  | ms | 10 | y2 | 10.7
| frames/s | 7 |  | frames/s | 7
| Fan type | Full |  | Fan type | Full
| Arc type | Half |  | Arc type | Half
kV | 122.89 | kV | kV | 121.1 | kV
HVL | 7.849 | mm Al | HVL | 7.857 | mm Al
Total filt | 9.8831 | mmAl TF | Total filt | 10.8568 | mmAl TF
Dose | 0.689 | mGy | Dose | 0.556 | mGy
Dose rate | 0.066 | mGy/s | Dose rate | 0.006 | mGy/s
frame/s | - | frame/s | frame/s | - | frame/s

## Thorax
Thorax |  |  | x1 | 24.7 | cm
| kvp | 125 | x2 | 3.3
| mA | 15 | y1 | 10.7
| ms | 20 | y2 | 10.7
| frames/s | 15
| Fan type | Half
| Arc type | Full
kV | 123.96 | kV
HVL | 7.953 | mm Al
Total filt | - | mmAl TF
Dose | 3.839 | mGy
Dose rate | 0.321 | mGy/s
frame/s | 14.7658 | frame/s

## Image gently
Image gently |  |  | x1 | 14 | rf low
| kvp | 80 | x2 | 14
| mA | 20 | y1 | 10.7
| ms | 10 | y2 | 10.7
| frames/s | 15
| Fan type | Full
| Arc type | Half
kV | 79.81 | kV
HVL | 5.737 | mm Al
Total filt | 10.9072 | mmAl TF
Dose | 0.631 | mGy
Dose rate | 0.058 | mGy/s
frame/s | 14.7813 | frame/s

## Collated data
Head | Mode | RF high | Pelvis | Mode | RF high | Pelvis large | Mode | RF high | Breast 360 | Mode | RF high | Breast 200 | Mode | RF high | Thorax | Mode | RF high | Image gently | Mode | RF low
| kvp | 100 |  | kvp | 125 |  | kvp | 140 |  | kvp | 125 |  | kvp | 125 |  | kvp | 125 |  | kvp | 80
| mA | 15 |  | mA | 60 |  | mA | 75 |  | mA | 10 |  | mA | 10 |  | mA | 15 |  | mA | 20
| ms | 20 |  | ms | 20 |  | ms | 25 |  | ms | 10 |  | ms | 10 |  | ms | 20 |  | ms | 10
| frames/s | 15 |  | frames/s | 15 |  | frames/s | 5 |  | frames/s | 7 |  | frames/s | 7 |  | frames/s | 15 |  | frames/s | 15
| Fan type | Full |  | Fan type | Half |  | Fan type | Half |  | Fan type | Full |  | Fan type | Full |  | Fan type | Half |  | Fan type | Full
| Arc type | Half |  | Arc type | Full |  | Arc type | Full |  | Arc type | Half |  | Arc type | Half |  | Arc type | Full |  | Arc type | Half
kV | 100.38 | kV | kV | 129.04 | kV | kV | 143.29 | kV | kV | 122.89 | kV | kV | 121.1 | kV | kV | 123.96 | kV | kV | 79.81 | kV
HVL | 6.887 | mm Al | HVL | 8.371 | mm Al | HVL | 8.942 | mm Al | HVL | 7.849 | mm Al | HVL | 7.857 | mm Al | HVL | 7.953 | mm Al | HVL | 5.737 | mm Al
Total filt | 10.829 | mmAl TF | Total filt | - | mmAl TF | Total filt | - | mmAl TF | Total filt | 9.8831 | mmAl TF | Total filt | 10.8568 | mmAl TF | Total filt | - | mmAl TF | Total filt | 10.9072 | mmAl TF
Dose | 1.392 | mGy | Dose | 9.28 | mGy | Dose | 6.008 | mGy | Dose | 0.689 | mGy | Dose | 0.556 | mGy | Dose | 3.839 | mGy | Dose | 0.631 | mGy
Dose rate | 0.171 | mGy/s | Dose rate | 1.11 | mGy/s | Dose rate | 0.75 | mGy/s | Dose rate | 0.066 | mGy/s | Dose rate | 0.006 | mGy/s | Dose rate | 0.321 | mGy/s | Dose rate | 0.058 | mGy/s
frame/s | 14.7489 | frame/s | frame/s | 14.7574 | frame/s | frame/s | 4.9669 | frame/s | frame/s | - | frame/s | frame/s | - | frame/s | frame/s | 14.7658 | frame/s | frame/s | 14.7813 | frame/s
Blades |  |  | Blades |  |  | Blades |  |  | Blades |  |  | Blades |  |  | Blades |  |  | Blades
x1 | 14 | cm | x1 | 24.7 | cm | x1 | 24.7 | cm | x1 | 14 | cm | x1 | 14 | cm | x1 | 24.7 | cm | x1 | 14 | cm
x2 | 14 |  | x2 | 3.3 |  | x2 | 3.3 |  | x2 | 14 |  | x2 | 14 |  | x2 | 3.3 |  | x2 | 14
y1 | 10.7 |  | y1 | 10.7 |  | y1 | 10.7 |  | y1 | 10.7 |  | y1 | 10.7 |  | y1 | 10.7 |  | y1 | 10.7
y2 | 10.7 |  | y2 | 10.7 |  | y2 | 10.7 |  | y2 | 10.7 |  | y2 | 10.7 |  | y2 | 10.7 |  | y2 | 10.7

## New tests August
|  | Head | Mode | RF high | [7 empty] | Pelvis | Mode | RF high |  | Breast 360 | Mode | RF high |  |  | Pelvis large | Mode | RF high |  | Breast 200 | Mode | RF high |  | Thorax | Mode | RF high |  | Image gently | Mode | RF low
|  |  | kvp | 100 | [8 empty] | kvp | 125 |  |  | kvp | 125 |  |  |  | kvp | 140 |  |  | kvp | 125 |  |  | kvp | 125 |  |  | kvp | 80
|  |  | mA | 15 | [8 empty] | mA | 60 |  |  | mA | 10 |  |  |  | mA | 75 |  |  | mA | 10 |  |  | mA | 15 |  |  | mA | 20
|  |  | ms | 20 | [8 empty] | ms | 20 |  |  | ms | 10 |  |  |  | ms | 25 |  |  | ms | 10 |  |  | ms | 20 |  |  | ms | 10
|  |  | frames/s | 15 | [8 empty] | frames/s | 15 |  |  | frames/s | 7 |  |  |  | frames/s | 5 |  |  | frames/s | 7 |  |  | frames/s | 15 |  |  | frames/s | 15
|  |  | Fan type | Full | [8 empty] | Fan type | Half |  |  | Fan type | Full |  |  |  | Fan type | Half |  |  | Fan type | Full |  |  | Fan type | Half |  |  | Fan type | Full
|  |  | Arc type | Half | [8 empty] | Arc type | Full |  |  | Arc type | Half |  |  |  | Arc type | Full |  |  | Arc type | Half |  |  | Arc type | Full |  |  | Arc type | Half
|  | INPLANE | [5 empty] | CROSSPLANE
|  | Couch (cm) | Dose rate uGy/s | Total Filtration |  |  |  | Couch (cm) | Dose rate uGy/s | Total Filtration |  | Couch (cm) | Dose rate uGy/s | HVL (TF not working) |  | Couch (cm) | Dose rate uGy/s | HVL Al | Total Filtration |  | Couch (cm) | Dose rate uGy/s | Total Filtration |  | Couch (cm) | Dose rate uGy/s | Total Filtration |  | Couch (cm) | Dose rate uGy/s | Total Filtration |  | Couch (cm) | Dose rate uGy/s | Total Filtration
|  |  |  | HVL in Aarons work?? | [8 empty] | B WALL |  |  |  | B WALL |  |  |  |  | B WALL
| 99.6 | 15 | SUP |  |  |  | 15.06 | 15 | B WALL |  | 15.06 | 15 |  |  | 15.06 | 15 |  |  |  | 15.06 | 15 |  |  |  | 15 |  |  |  | 15 |  |  |  | 15
| 99.1 | 14.5 |  |  |  |  | 14.56 | 14.5 |  |  | 14.56 | 14.5 |  |  | 14.56 | 14.5 |  |  |  | 14.56 | 14.5 |  |  |  | 14.5 |  |  |  | 14.5 |  |  |  | 14.5
| 98.6 | 14 |  |  |  |  | 14.06 | 14 |  |  | 14.06 | 14 |  |  | 14.06 | 14 |  |  |  | 14.06 | 14 |  |  |  | 14 |  |  |  | 14 |  |  |  | 14
| 98.1 | 13.5 |  |  |  |  | 13.56 | 13.5 |  |  | 13.56 | 13.5 |  |  | 13.56 | 13.5 |  |  |  | 13.56 | 13.5 |  |  |  | 13.5 |  |  |  | 13.5 |  |  |  | 13.5
| 97.6 | 13 |  |  |  |  | 13.06 | 13 |  |  | 13.06 | 13 |  |  | 13.06 | 13 |  |  |  | 13.06 | 13 |  |  |  | 13 |  |  |  | 13 |  |  |  | 13
| 97.1 | 12.5 |  |  |  |  | 12.56 | 12.5 |  |  | 12.56 | 12.5 |  |  | 12.56 | 12.5 |  |  |  | 12.56 | 12.5 |  |  |  | 12.5 |  |  |  | 12.5 |  |  |  | 12.5
| 96.6 | 12 |  |  |  |  | 12.06 | 12 |  |  | 12.06 | 12 |  |  | 12.06 | 12 |  |  |  | 12.06 | 12 |  |  |  | 12 |  |  |  | 12 |  |  |  | 12
| 96.1 | 11.5 |  |  |  |  | 11.56 | 11.5 |  |  | 11.56 | 11.5 |  |  | 11.56 | 11.5 |  |  |  | 11.56 | 11.5 |  |  |  | 11.5 |  |  |  | 11.5 |  |  |  | 11.5
| 95.6 | 11 |  |  |  |  | 11.06 | 11 | colide with gantry head can't go further |  | 11.06 | 11 |  |  | 11.06 | 11 |  |  |  | 11.06 | 11 |  |  |  | 11 |  |  |  | 11 |  |  |  | 11
| 95.1 | 10.5 |  |  |  |  | 10.56 | 10.5 |  |  | 10.56 | 10.5 |  |  | 10.56 | 10.5 | 186 | 10.78 |  | 10.56 | 10.5 | 1687 | 8.477 |  | 10.5 |  |  |  | 10.5 |  |  |  | 10.5
| 94.6 | 10 | too low for rdg |  |  |  | 10.06 | 10 | 67.61 | 30.54 | 10.06 | 10 | 1008 | 7.891 | 10.06 | 10 | 196.3 | 10.67 |  | 10.06 | 10 | 1695 | 8.551 |  | 10 |  |  |  | 10 |  |  |  | 10
| 94.1 | 9.5 | 251.7 |  |  |  | 9.56 | 9.5 |  |  | 9.56 | 9.5 |  |  | 9.56 | 9.5 |  |  |  | 9.56 | 9.5 |  |  |  | 9.5 |  |  |  | 9.5 |  |  |  | 9.5
| 93.6 | 9 | 202.8 | 9 |  |  | 9.06 | 9 | 79.31 | 22.48 | 9.06 | 9 |  |  | 9.06 | 9 | 221.6 | 10.681 |  | 9.06 | 9 |  |  |  | 9 |  |  |  | 9 |  |  |  | 9
| 93.35 | 8.75 | 367.3 |  |  |  | 8.81 | 8.75 |  |  | 0.06 |  |  |  | 0.06 |  |  |  |  | 0.06
#VALUE! | 93.1 | 8.5 | 418 | 17.25 |  |  | 8.56 | 8.5 | 93.62 | 19.6 | 8.56 | 8.5 |  |  | 8.56 | 8.5 |  |  |  | 8.56 | 8.5 |  |  |  | 8.5 |  |  |  | 8.5 |  |  |  | 8.5
| 92.85 | 8.25 | 434 | 12.28 |  |  | 8.31 | 8.25 |  |  | 0.06 |  |  |  | 0.06 |  |  |  |  | 0.06
| 92.6 | 8 | 457.9 | 9.257 |  |  | 8.06 | 8 | 124 | 25.85 | 8.06 | 8 | 1007 | 8.17 | 8.06 | 8 | 329.4 | 10.101 |  | 8.06 | 8 | 1704 | 8.592 |  | 8 |  |  |  | 8 |  |  |  | 8
| 92.1 | 7.5 |  |  |  |  | 7.56 | 7.5 |  |  | 7.56 | 7.5 |  |  | 7.56 | 7.5 |  |  |  | 7.56 | 7.5 |  |  |  | 7.5 |  |  |  | 7.5 |  |  |  | 7.5
| 91.6 | 7 | 458.3 | 10.09 |  |  | 7.06 | 7 | 200.7 | 16.11 | 7.06 | 7 |  |  | 7.06 | 7 | 501.2 | 9.641 |  | 7.06 | 7 |  |  |  | 7 |  |  |  | 7 |  |  |  | 7
| 91.1 | 6.5 |  |  |  |  | 6.56 | 6.5 |  |  | 6.56 | 6.5 |  |  | 6.56 | 6.5 |  |  |  | 6.56 | 6.5 |  |  |  | 6.5 |  |  |  | 6.5 |  |  |  | 6.5
| 90.6 | 6 | 465.6 | 9.584 |  |  | 6.06 | 6 | 256.6 | 16.75 | 6.06 | 6 | 1029 | 7.905 | 6.06 | 6 | 628.9 | 9.146 |  | 6.06 | 6 | 1720 | 8.638 |  | 6 |  |  |  | 6 |  |  |  | 6
| 90.1 | 5.5 |  |  |  |  | 5.56 | 5.5 |  |  | 5.56 | 5.5 |  |  | 5.56 | 5.5 |  |  |  | 5.56 | 5.5 |  |  |  | 5.5 |  |  |  | 5.5 |  |  |  | 5.5
| 89.6 | 5 | 462.1 | 10.82 |  |  | 5.06 | 5 | 312.7 | 14.16 | 5.06 | 5 |  |  | 5.06 | 5 |  |  |  | 5.06 | 5 |  |  |  | 5 |  |  |  | 5 |  |  |  | 5
| 89.1 | 4.5 |  |  |  |  | 4.56 | 4.5 |  |  | 4.56 | 4.5 |  |  | 4.56 | 4.5 |  |  |  | 4.56 | 4.5 |  |  |  | 4.5 |  |  |  | 4.5 |  |  |  | 4.5
G=90, souce at top | 88.6 | 4 | 453.4 | 9.257 |  |  | 4.06 | 4 | 362 | 13.04 | 4.06 | 4 | 1039 | 8.103 | 4.06 | 4 | 835.6 | 8.434 |  | 4.06 | 4 | 1730 | 8.582 |  | 4 |  |  |  | 4 |  |  |  | 4
couch moved in LONG | 88.1 | 3.5 |  |  |  |  | 3.56 | 3.5 |  |  | 3.56 | 3.5 |  |  | 3.56 | 3.5 |  |  |  | 3.56 | 3.5 |  |  |  | 3.5 |  |  |  | 3.5 |  |  |  | 3.5
detector long axis on couch long axis | 87.6 | 3 | 466 | 9.564 |  |  | 3.06 | 3 | 399.2 | 11.71 | 3.06 | 3 |  |  | 3.06 | 3 |  |  |  | 3.06 | 3 |  |  |  | 3 |  |  |  | 3 |  |  |  | 3
| 87.1 | 2.5 |  |  |  |  | 2.56 | 2.5 |  |  | 2.56 | 2.5 |  |  | 2.56 | 2.5 |  |  |  | 2.56 | 2.5 |  |  |  | 2.5 |  |  |  | 2.5 |  |  |  | 2.5
| 86.6 | 2 | 470 | 8.52 |  |  | 2.06 | 2 | 430 | 11.61 | 2.06 | 2 | 1024 | 8.044 | 2.06 | 2 | 961.8 | 8.049 |  | 2.06 | 2 | 1727 | 8.627 |  | 2 |  |  |  | 2 |  |  |  | 2
| 86.1 | 1.5 |  |  |  |  | 1.56 | 1.5 |  |  | 1.56 | 1.5 |  |  | 1.56 | 1.5 |  |  |  | 1.56 | 1.5 |  |  |  | 1.5 |  |  |  | 1.5 |  |  |  | 1.5
Couch | 85.6 | 1 | 471.1 | 8.474 |  | Couch | 1.06 | 1 | 443.3 | 10.77 | 1.06 | 1 |  |  | 1.06 | 1 |  |  |  | 1.06 | 1 |  |  | Couch | 1 |  |  | Couch | 1 |  |  | Couch | 1
coordinates | 85.1 | 0.5 |  |  |  | coordinates | 0.56 | 0.5 |  |  | 0.56 | 0.5 |  |  | 0.56 | 0.5 |  |  |  | 0.56 | 0.5 |  |  | coordinates | 0.5 |  |  | coordinates | 0.5 |  |  | coordinates | 0.5
LONG | 84.6 | 0 | 469 | 8.034 |  | LAT | 0.06 | 0 | 447.2 | 11.57 | 0.06 | 0 | 1018 | 8.06 | 0.06 | 0 | 1002 | 8.045 |  | 0.06 | 0 | 1706 | 8.539 | 0 | 0 |  |  | 0 | 0 |  |  | 0 | 0
| 84.1 | -0.5 |  |  |  |  | 999.56 | -0.5 |  |  | 999.56 | -0.5 |  |  | 999.56 | -0.5 |  |  |  | 999.56 | -0.5 |  |  | 999.5 | -0.5 |  |  | 999.5 | -0.5 |  |  | 999.5 | -0.5
| 83.6 | -1 |  |  |  |  | 999.06 | -1 | 434.4 | 11.18 | 999.06 | -1 |  |  | 999.06 | -1 |  |  |  | 999.06 | -1 |  |  | 999 | -1 |  |  | 999 | -1 |  |  | 999 | -1
| 83.1 | -1.5 |  |  |  |  | 998.56 | -1.5 |  |  | 998.56 | -1.5 |  |  | 998.56 | -1.5 |  |  |  | 998.56 | -1.5 |  |  | 998.5 | -1.5 |  |  | 998.5 | -1.5 |  |  | 998.5 | -1.5
| 82.6 | -2 | 445.6 | 10.37 |  |  | 998.06 | -2 |  |  | 998.06 | -2 | 971.5 | 8.146 | 998.06 | -2 | 935.1 | 8.34 |  | 998.06 | -2 | 1639 | 8.645 | 998 | -2 |  |  | 998 | -2 |  |  | 998 | -2
| 82.1 | -2.5 |  |  |  |  | 997.56 | -2.5 |  |  | 997.56 | -2.5 |  |  | 997.56 | -2.5 |  |  |  | 997.56 | -2.5 |  |  | 997.5 | -2.5 |  |  | 997.5 | -2.5 |  |  | 997.5 | -2.5
| 81.6 | -3 |  |  |  |  | 997.06 | -3 | 379 | 11.85 | 997.06 | -3 |  |  | 997.06 | -3 |  |  |  | 997.06 | -3 |  |  | 997 | -3 |  |  | 997 | -3 |  |  | 997 | -3
| 81.1 | -3.5 |  |  |  |  | 996.56 | -3.5 |  |  | 996.56 | -3.5 |  |  | 996.56 | -3.5 |  |  |  | 996.56 | -3.5 |  |  | 996.5 | -3.5 |  |  | 996.5 | -3.5 |  |  | 996.5 | -3.5
| 80.6 | -4 | 438.8 | 10.05 |  |  | 996.06 | -4 |  |  | 996.06 | -4 | 878.6 | 8.353 | 996.06 | -4 | 771.7 | 8.639 |  | 996.06 | -4 | 1471 | 8.993 | 996 | -4 |  |  | 996 | -4 |  |  | 996 | -4
| 80.1 | -4.5 |  |  |  |  | 995.56 | -4.5 |  |  | 995.56 | -4.5 |  |  | 995.56 | -4.5 |  |  |  | 995.56 | -4.5 |  |  | 995.5 | -4.5 |  |  | 995.5 | -4.5 |  |  | 995.5 | -4.5
| 79.6 | -5 |  |  |  |  | 995.06 | -5 | 273 | 18.64 | 995.06 | -5 |  |  | 995.06 | -5 |  |  |  | 995.06 | -5 |  |  | 995 | -5 |  |  | 995 | -5 |  |  | 995 | -5
| 79.1 | -5.5 |  |  |  |  | 994.56 | -5.5 |  |  | 994.56 | -5.5 |  |  | 994.56 | -5.5 |  |  |  | 994.56 | -5.5 |  |  | 994.5 | -5.5 |  |  | 994.5 | -5.5 |  |  | 994.5 | -5.5
| 78.6 | -6 | 430.4 | 10.13 |  |  | 994.06 | -6 |  |  | 994.06 | -6 | 739.7 | 8.809 | 994.06 | -6 | 543.1 | 9.607 |  | 994.06 | -6 | 1249 | 9.431 | 994 | -6 |  |  | 994 | -6 |  |  | 994 | -6
| 78.1 | -6.5 |  |  |  |  | 993.56 | -6.5 |  |  | 993.56 | -6.5 |  |  | 993.56 | -6.5 |  |  |  | 993.56 | -6.5 |  |  | 993.5 | -6.5 |  |  | 993.5 | -6.5 |  |  | 993.5 | -6.5
| 77.6 | -7 |  |  |  |  | 993.06 | -7 | 153.1 | 19.55 | 993.06 | -7 |  |  | 993.06 | -7 | 389.9 | 10.094 |  | 993.06 | -7 |  |  | 993 | -7 |  |  | 993 | -7 |  |  | 993 | -7
| 77.1 | -7.5 |  |  |  |  | 992.56 | -7.5 |  |  | 992.56 | -7.5 |  |  | 992.56 | -7.5 |  |  |  | 992.56 | -7.5 |  |  | 992.5 | -7.5 |  |  | 992.5 | -7.5 |  |  | 992.5 | -7.5
| 76.6 | -8 | 422.1 | 10.13 |  |  | 992.06 | -8 | 86.19 | 19.49 | 992.06 | -8 | 557.6 | 9.204 | 992.06 | -8 | 235.2 | 10.756 |  | 992.06 | -8 | 963.7 | 9.893 | 992 | -8 |  |  | 992 | -8 |  |  | 992 | -8
| 76.1 | -8.5 | 421 | 10.35 |  |  | 991.56 | -8.5 |  |  | 991.56 | -8.5 |  |  | 991.56 | -8.5 |  |  |  | 991.56 | -8.5 |  |  | 991.5 | -8.5 |  |  | 991.5 | -8.5 |  |  | 991.5 | -8.5
| 75.6 | -9 | 418 | 10.21 |  |  | 991.06 | -9 | 70.22 | 23.17 | 991.06 | -9 |  |  | 991.06 | -9 | 197.3 | 10.867 | 25.56 | 991.06 | -9 | 802.1 | 10.373 | 991 | -9 |  |  | 991 | -9 |  |  | 991 | -9
| 75.1 | -9.5 |  |  |  |  | 990.56 | -9.5 |  |  | 990.56 | -9.5 |  |  | 990.56 | -9.5 |  |  |  | 990.56 | -9.5 |  |  | 990.5 | -9.5 |  |  | 990.5 | -9.5 |  |  | 990.5 | -9.5
| 74.6 | -10 | 424.8 |  |  |  | 990.06 | -10 | 64.1 | 24.14 | 990.06 | -10 | 340.6 | 9.958 | 990.06 | -10 | 183.6 | 11.16 |  | 990.06 | -10 | 605.5 | 10.759 | 990 | -10 |  |  | 990 | -10 |  |  | 990 | -10
| 74.1 | -10.5 |  |  |  |  | 989.56 | -10.5 |  |  | 989.56 | -10.5 |  |  | 989.56 | -10.5 |  |  |  | 989.56 | -10.5 |  |  | 989.5 | -10.5 |  |  | 989.5 | -10.5 |  |  | 989.5 | -10.5
| 73.6 | -11 |  |  |  |  | 989.06 | -11 | 60.83 | 25.42 | 989.06 | -11 |  |  | 989.06 | -11 | 172.8 | 10.629 |  | 989.06 | -11 | 420.5 | 11.279 | 989 | -11 |  |  | 989 | -11 |  |  | 989 | -11
| 73.1 | -11.5 |  |  |  |  | 988.56 | -11.5 |  |  | 988.56 | -11.5 |  |  | 988.56 | -11.5 |  |  |  | 988.56 | -11.5 |  |  | 988.5 | -11.5 |  |  | 988.5 | -11.5 |  |  | 988.5 | -11.5
| 72.6 | -12 |  |  |  |  | 988.06 | -12 | 58.72 | 26.5 | 988.06 | -12 | 195.8 | 10.571 | 988.06 | -12 | 163.1 | 11.227 | 31.46 | 988.06 | -12 | 352.9 | 11.333 | 988 | -12 |  |  | 988 | -12 |  |  | 988 | -12
| 72.1 | -12.5 |  |  |  |  | 987.56 | -12.5 |  |  | 987.56 | -12.5 |  |  | 987.56 | -12.5 |  |  |  | 987.56 | -12.5 |  |  | 987.5 | -12.5 |  |  | 987.5 | -12.5 |  |  | 987.5 | -12.5
| 71.6 | -13 |  |  |  |  | 987.06 | -13 | 54.67 | 21.51 | 987.06 | -13 | 169.9 | 10.06 | 987.06 | -13 | 154.7 | 10.95 |  | 987.06 | -13 | 300.6 | 11.344 | 987 | -13 |  |  | 987 | -13 |  |  | 987 | -13
| 71.1 | -13.5 |  |  |  |  | 986.56 | -13.5 |  |  | 986.56 | -13.5 |  |  | 986.56 | -13.5 |  |  |  | 986.56 | -13.5 |  |  | 986.5 | -13.5 |  |  | 986.5 | -13.5 |  |  | 986.5 | -13.5
| 70.6 | -14 |  |  |  |  | 986.06 | -14 | 13.51 | 14.92 | 986.06 | -14 | 42.3 | 8.589 | 986.06 | -14 | 89.36 | 10.807 |  | 986.06 | -14 | 72.78 | 9.751 | 986 | -14 |  |  | 986 | -14 |  |  | 986 | -14
| 70.1 | -14.5 |  |  |  |  | 985.56 | -14.5 |  |  | 985.56 | -14.5 |  |  | 985.56 | -14.5 |  |  |  | 985.56 | -14.5 |  |  | 985.5 | -14.5 |  |  | 985.5 | -14.5 |  |  | 985.5 | -14.5
| 69.6 | -15 | INF |  |  |  | 985.06 | -15 | 6.345 | 8.776 | 985.06 | -15 | 22.68 | 8.026 | 985.06 | -15 | 18.62 | 8.703 | 19.18 | 985.06 | -15 | 39.53 | 8.628 | 985 | -15 |  |  | 985 | -15 |  |  | 985 | -15
[9 empty] | A WALL |  |  |  | A WALL |  |  |  | A WALL
[5 empty] | NOTE ms=30 for these measurements as device didn't work for ms=15 | [6 empty] | NOTE: used 20mA instead of 60 to keep readings in range
NOTE: used 20mA instead of 10 as service mode wouldn't let us do planar that low
Check consistency of acquisitions (each measuremetn fro ma individal DynamicGain shot. In XI tab. Set appropriate Filter Shape (HalfFan/FullFan) in Position Units tab -->KVs coll-->clinical axes
| mGy | 19.7 | 19.51 | 19.69 | [6 empty] | 19.6333 | 0.10693 | 0.5446%
| mGy/s | 550.5 | 546.9 | 552 | [6 empty] | 549.8 | 2.6211 | 0.4767%
| filtration | 7.328 | 7.697 | 7.452 | [6 empty] | 7.4923 | 0.18778 | 2.506%
Note on D45: Izabela Junis (Western Sydney LHD): 44.72 on repeat.. Seemed to have reset between two halves of profile
Note on J45: Izabela Junis (Western Sydney LHD): 450.5 repeat meas at end of profile set

## Sheet1
Breast 360 | Mode | RF high
| kvp | 125
| mA | 10
| ms | 10
| frames/s | 7
| Fan type | Full
| Arc type | Half