## Sheet1
Current | [12 empty] | Only trust dose factor for Stds, otherwise it miught just be copied over not measured
[12 empty] | deg/s | mGy/100mAs | surrogate for half/full fan? |  | deg | s |  | cGy
LA3 | LA4 | LA5 | CST | OBK | Mode Name | Std/Custom | Clinical? | kV | mA/frame | ms/frame | frames/s | Gantry speed | Dose Factor | CTDI Phantom | No Projections | Proj Increment | Acqusition time | Exposure | CTDI | Trajectory | Start angle | Prescan darkfield | Algo | Filter | Matrix size | Slice thick | Ring Suppresion | Fan | Column1 | Notes
Yes | Yes | Yes | Yes | Yes | 4D spotlight | Standard | Yes | 125 | 40 | 20 | 7 | 3 | 1.64 | 320mm | 467 | 0.428 | 67 | 371.6 | 0.61 | Half | Auto | No | 4D | Auto | 384 | 1.5 | No | Full (SFOV)
Yes | Yes | Yes | Yes | Yes | 4D thorax | Standard | Yes | 125 | 40 | 20 | 7 | 3 | 1.48 | 320mm | 840 | 0.429 | 120 | 672 | 1 | Full |  | No | 4D | Auto | 384 | 1.5 | No | Half (LFOV)
Yes | Yes | Yes | Yes | Yes | Breast 360 | Custom | Yes | 125 | 10 | 10 | 7 | 6 | 1.48 | 320mm | 420 | 0.857 | 60 | 42 | 0.06 | Full |  | No | Default | Auto | 512 | 2 | Medium | Half (SFOV) |  | IS HALF CORRECT????
Yes | Yes | Yes | Yes | Yes | Head | Standard | Yes | 100 | 15 | 20 | 15 | 6 | 2.11 | 160mm | 500 | 0.4 | 33 | 150 | 0.32 | Half | Auto | No | Default | Auto | 512 | 2 | Medium | Full (SFOV)
Yes |  |  |  |  | Head Large FOV | Custom | Yes | 100 | 10 | 20 | 15 | 6 | 2.11 | 160mm | 900 | 0.4 | 60 | 180 | 0.38 | Full |  | No | Default | Auto | 512 | 2 | Medium | Half (LFOV)
|  | Yes |  |  | Head_low_dose | Custom | Yes | 100 | 10 | 10 | 7 | 6 | 1.95 | 160mm | 420 | 0.857 | 30 | 42 | 0.08 | Full |  | No | Default | Auto | 512 | 1 | Medium | Full (SFOV) |  | ?? Attempt at Paed head?? Or SRS (1mm recon)
| Yes | Yes |  |  | HN half fan | Custom | Yes | 125 | 10 | 10 | 15 | 6 | 1.32 | 320mm | 900 | 0.4 | 60 | 90 | 0.12 | Full |  | Yes | Default | Auto | 512 | 2 | Medium | Half (LFOV) |  | same as Head Large FOV?? Just CTDI entered incorrectly
Yes | Yes | Yes | Yes | Yes | Image Gently | Standard | Yes | 80 | 20 | 10 | 15 | 6 | 0.94 | 160mm | 500 | 0.4 | 33 | 100 | 0.09 | Half | Auto | No | Default | Auto | 512 | 2 | Medium | Full (SFOV)
| Yes |  |  |  | KIM | Custom | Yes | 125 | 60 | 25 | 15 | 6 | 1.64 | 320mm | 900 | 0.4 | 60 | 1350 | 2.21 | Half |  | No | Default | Smooth | 512 | 2 | Medium | Half (LFOV)
Yes |  |  |  |  | KIM_LARK | Custom | Yes | 125 | 60 | 25 | 7 | 4 | 1.64 | 320mm | 350 | 0.571 | 50 | 525 | 0.86 | Half | Auto | Yes | Default | Smooth | 512 | 2 | Medium |  |  | Prescan only on LA5 not LA4
Yes | Yes | Yes | Yes | Yes | Pelvis | Standard | Yes | 125 | 60 | 20 | 15 | 6 | 1.48 | 320mm | 900 | 0.4 | 60 | 1080 | 1.6 | Full | Auto | No | Default | Auto | 512 | 2 | Medium | Half (LFOV)
Yes | Yes | Yes | Yes | Yes | Pelvis Obese | Standard | Yes | 140 | 75 | 25 | 15 | 6 | 2.18 | 320mm | 900 | 0.4 | 60 | 1687 | 3.68 | Full | Auto | No | Default | Auto | 512 | 2 | Medium | Half (LFOV)
|  | Yes |  |  | physics_breast | Custom | Yes | 100 | 15 | 20 | 15 | 6 | 1.95 | 160mm | 500 | 0.4 | 33 | 150 | 0.29 | Half | Manual (180--20) | No
Yes | Yes |  | Yes | Yes | SBRT Lung | Custom | Yes | 125 | 15 | 20 | 15 | 4.5 | 1.32 | 160mm | 667 | 0.3 | 44 | 200 | 0.26 | Half | Auto | No | Default | Auto | 512 | 2 | Medium | Full (SFOV)
Yes | Yes |  | Yes | Yes | SBRT spine midDose | Custom | Yes | 125 | 40 | 10 | 15 | 6 | 1.48 | 320mm | 900 | 0.4 | 60 | 360 | 0.53 | Full | Auto | No | Default | Auto | 512 | 2 | Medium | Full (SFOV)
Yes | Yes |  | Yes | Yes | SBRT_Spine_Lowdose | Custom | Yes | 125 | 20 | 10 | 15 | 6 | 1.48 | 320mm | 900 | 0.4 | 60 | 180 | 0.27 | Full | Auto | No | Default | Auto | 512 | 2 | Medium | Full (SFOV)
Yes | Yes | Yes | Yes | Yes | Spotlight | Standard | Yes | 125 | 60 | 25 | 15 | 6 | 1.64 | 320mm | 500 | 0.4 | 33 | 750 | 1.23 | Half | Auto | No | Default | Smooth | 512 | 2 | Medium | Full (SFOV)
Yes | Yes | Yes | Yes | Yes | Thorax | Standard | Yes | 125 | 15 | 20 | 15 | 6 | 1.48 | 320mm | 900 | 0.4 | 60 | 270 | 0.4 | Full |  | No | Default | Smooth | 512 | 2 | Medium | Half (LFOV)
| Yes |  |  |  | 4D SBRT Patient | Custom | Non-clinical | 125 | 40 | 20 | 7 | 3 | 1.32 | 320mm | 840 | 0.429 | 120 | 672 | 0.89 | Full |  | No | 4D | Auto | 384 | 1.5 | Medium | Full (SFOV)
|  | Yes |  |  | Rabbit |  | Non-clinical | [22 empty] | From old JB project, delete
|  | Yes |  |  | Rabbit_Slow |  | Non-clinical | [22 empty] | From old JB project, delete
| Yes |  |  |  | Pelvis_VHMP1 | Custom | Yes | [22 empty] | From old JB project, delete
| Yes |  |  |  | Pelvis_VHMP2 | Custom | Yes | [22 empty] | From old JB project, delete
[16 empty] | mA | ms | fps | G speed | G range | Exposure | #projections | Time
[16 empty] | 15 | 20 | 11 | 6 | 360 | 198 | 660 | 60
Planned List (uniform across all machines)
| Standards |  | Customs |  | Note
| 4D Spotlight |  |  |  | 200deg half-fan 4D
| 4D Thorax |  |  |  | 360deg full-fan 4D
|  |  | Abdomen |  | ?Copy of pelvis with 75% mAs? | [7 empty] | Copy of Thorax, mA=40
|  |  | Breast 360 |  | 360deg full-fan, requires couch centring
|  |  | ?Breast left 200 |  | ideally would not require couch-centring
|  |  | ?Breast right 200 |  | ideally would not require couch-centring
|  |  | Extremity |  | ?needed… use copy of Head | [7 empty] | Copy of Head
| Head |  |  |  | 200deg full-fan, 2mm recon .. Could modify to always shoot from behind lens eg, G=180-->20???
|  |  | Head SRS |  | 200deg full-fan, 1mm recon in z | [7 empty] | Copy of Head, 1mm recon, full rotation, 20mA, 10ms (X=180)
|  |  | Head Paediatric |  | copy of Head with half mAs
|  |  | Head and Shoulders |  | ?use copy of Head with LFOV (like Head Large FOV) or copy of Thorax… ultimately LFOV 360deg half-fan | [7 empty] | Copy of thorax
| Image Gently |  |  |  | ? Copy and rename Paediatric body?
|  |  | Paediatric Body | [9 empty] | Copy of IG
|  |  | Paediatric Head |  | copy of Head with half mAs | [7 empty] | Copy of Head, 80kV, 10mA 10ms. Dose factor 0.94 per IG
| Pelvis |  |  |  | 360deg, half-fan, high exposure
Pelvis-Obese
|  |  | SBRT lung |  | copy of thorax with half rotation and slower gantry rotation… quicker than throax but some extra motion blur for ITV matching?
|  |  | SBRT spine mid-dose |  | initial CBCT, 360deg half-fan | [7 empty] | 1mm res
|  |  | SBRT spine low-dose |  | subsequent CBCT, same scan half mAs… ?could be 200deg scan to speed up? | [7 empty] | 1mm res
| Spotlight |  |  |  | small FOV, high dose for pelvis (hard to directly compare but about 75% dose or Pelvis)
| Thorax |  |  |  | 360deg, half-fan, lower exposure
KIM
KIM Lark
General questions:
Can gantry start/stop be optimised rathe than using auto always?
Are extremity and abo needed? Any others missing?
Actions
Delete VHMP settings and rabbit settings from LA4 and LA5 respectively
Rename and standardise HN LFOVand HN half fan
Decide on, and add new customs (Abdo, Extremity, Breast 200's, Head SRS, Head Paed)