## Original
| Current | 2023-08-22
[7 empty] | Acquistion | [15 empty] | Reconstruction
[19 empty] | deg | s | mAs | cGy
Mode Name | LA3 | LA4 | LA5 | CST | OBK | Clinical | kV | mA/frame | ms/frame | frames/s | deg/s | CTDI Phantom | mGy/100mAs | Prescan darkfield | Trajectory | Start angle | Fan | No Projections | Proj Increment | Acqusition time | Exposure | CTDI | Default 3D Method | Default 4D Method | Algo | Filter | Matrix size | Slice thick | No. Phases | Phase Bin Width
4D Spotlight | Yes |  |  |  |  | Non-clinical | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | No | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  | Yes |  |  | Non-clinical | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | No | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  |  | Yes |  | Non-clinical | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | No | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
[5 empty] | Yes | Non-clinical | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | No | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
4D thorax | Yes |  |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  | Yes |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
[5 empty] | Yes | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
Abdomen | No
|  | Yes |  |  |  | Yes | 125 | 30 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.799 | Standard | - | - | - | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 10 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | Standard | - |  |  | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Default | - | - | - | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Standard | - | - | - | 512 | 2 | - | -
Breast 360 | Yes |  |  |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes (BH) |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes (BH) | Non-clinical | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Standard | Auto | 512 | 2 | - | -
Breast Left 200 | No
No
No
|  |  |  | Yes |  | Yes | 125 | 10 | 10 | 15 | 6 | 160mm | 1.32 | Yes | Half | Manual | Full | 500 | 0.4 | 33 | 50 | 0.066 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes (BH) |  | Yes | 100 | 10 | 10 | 15 | 6 | 160mm | 2.11 | Yes | Half | Manual | Full | 500 | 0.4 | 33 | 50 | 0.106 | Default | - | Default | Auto | 512 | 2 | - | -
No
Head | Yes |  |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
Head SRS | Yes |  |  |  |  | Yes | 100 | 30 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | None | - | Default | Auto | 512 | 1 | - | -
No
|  |  | Yes |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 270 | 0.57 | None | - | Default | Auto | 512 | 1 | - | -
|  |  |  | Yes |  | Yes | 100 | 20 | 10 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 180 | 0.38 | Default | - | Default | Auto | 512 | 1 | - | -
[5 empty] | Yes | Yes | 100 | 20 | 10 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 180 | 0.38 | HighRes | - | HighRes | Autp | 512 | ? | - | -
H&N half fan | No
|  | Yes |  |  |  | Yes | 125 | 10 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Standard | Auto | 512 | 2 | - | -
No
No
No
Head and Shoulders | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
Image Gently | Yes |  |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis | Yes |  |  |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Large (Pelvis Obese) | Yes |  |  |  |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
No
|  |  |  | Yes |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Spotlight (Spotlight) | Yes (double up) |  |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Yes (Double up) |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  |  | Yes (Double up) |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
[5 empty] | Yes (spotlight) | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
Short Thorax | Yes |  |  |  |  | Yes | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Yes |  |  |  | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax Spotlight (SBRT Lung) | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 160mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 160mm | 1.32 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 4.5 | 160mm | 1.32 | No | Half | Auto | Full (SFOV) | 667 | 0.3 | 44 | 200 | 0.264 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes (SBRT) | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Standard | Auto | 512 | 2 | - | -
Extremity Spotlight (extremity) | Yes |  |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
SBRT Spine MidDose | Yes |  |  |  |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
No
|  |  | Yes |  |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes (MedDose) |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.32 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.475 | Dafault | - | Default | Auto | 512 | 1 | - | -
[5 empty] | Yes (MedDose) | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
SBRT Spine LowDose | Yes |  |  |  |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
No
|  |  | Yes |  |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | No | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.32 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.238 | Dafault | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
Paediatric Head | Yes |  |  |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Default | Auto | 512 | 1 | - | -
|  | Yes |  |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Standard | Auto | 512 | 1 | - | -
|  |  |  | Yes |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | Default | - | Default | Auto | 512 | 1 | - | -
No
Paediatric Body | Yes |  |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | No | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
No
No
KIM LARK Large | Yes |  |  |  |  | Yes | 140 | 80 | 25 | 15 | 6 | 320mm | 2 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 1000 | 2 | None | - | Default | Smooth | 512 | 2 | - | -
No
No
No
No
KIM LARK Med | Yes |  |  |  |  | Yes | 125 | 68 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 850 | 1.394 | None | - | Default | Smooth | 512 | 2 | - | -
No
|  |  | Yes |  |  | Yes | 125 | 68 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 850 | 1.394 | None | - | Default | Smooth | 512 | 2 | - | -
No
No
KIM LARK Small | Yes |  |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Default | Smooth | 512 | 2 | - | -
No
|  |  | Yes |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Default | Smooth | 512 | 2 | - | -
No
No

## Updated (Change as actioned)
| Current | 2023-08-22
[7 empty] | Acquistion | [15 empty] | Reconstruction
[19 empty] | deg | s | mAs | cGy |  |  | ??
Mode Name | LA3 | LA4 | LA5 | CST | OBK | Clinical | kV | mA/frame | ms/frame | frames/s | deg/s | CTDI Phantom | mGy/100mAs | Prescan darkfield | Trajectory | Start angle | Fan | No Projections | Proj Increment | Acqusition time | Exposure | CTDI | Default 3D Method | Default 4D Method | Algo | Filter | Matrix size | Slice thick | No. Phases | Phase Bin Width
4D Spotlight | Yes |  |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  | Yes |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
[5 empty] | Yes | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
4D thorax | Yes |  |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  | Yes |  |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
[5 empty] | Yes | Yes | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
Abdomen | Yes |  |  |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Default | - | - | - | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Standard | - | - | - | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Standard | - |  |  | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Default | - | - | - | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Standard | - | - | - | 512 | 2 | - | -
Breast 360 | Yes |  |  |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | (BH deleted) |  | Non-clinical
[5 empty] | Yes | Yes | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | (BH deleted) | Non-clinical
Breast Left 200 | No
No
No
|  |  |  | Deleted | [22 empty] | 512 | 2 | - | -
|  |  |  | (BH deleted) | [22 empty] | 512 | 2 | - | -
No
Head | Yes |  |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | 10 | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
Head SRS | Yes |  |  |  |  | Yes | 100 | 30 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | None | - | Default | Auto | 512 | 1 | - | -
|  | Yes |  |  |  | Yes | 100 | 30 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | None | - | HighRes | Auto | 512 | 1 | - | -
|  |  | Yes |  |  | Yes | 100 | 30 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | None | - | Default | Auto | 512 | 1 | - | -
|  |  |  | Yes |  | Yes | 100 | 30 | 10 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | HighRes | - | HighRes | Auto | 512 | 1 | - | -
[5 empty] | Yes | Yes | 100 | 30 | 10 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | HighRes | - | HighRes | Auto | 512 | ? | - | -
H&N half fan | No
|  | No |  |  |  | Yes | 125 | 10 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Standard | Auto | 512 | 2 | - | -
No
No
No
Head and Shoulders | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
Image Gently | Yes |  |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis | Yes |  |  |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Large (Pelvis Obese) | Yes |  |  |  |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Spotlight (Spotlight) | Yes (double up) |  |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  |  | Yes ('Spotlight' now non-clinical) |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
[5 empty] | Yes (copied from 'Spotlight', which is now non-clinical) | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | z | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Auto | 512 | 2 | - | -
Short Thorax | Disabled |  |  |  |  | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  | Disabled |  |  |  | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Disabled |  |  | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  |  | Yes |  | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
[5 empty] | Yes | Non-clinical | 125 | 30 | 20 | 15 | 6 | 320mm | 1.64 | No | Limited 140 | Auto | Full (SFOV) | 350 | 0.4 | 23 | 210 | 0.344 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax Spotlight (SBRT Lung) | Yes |  |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes (rename to 'Thoax Spotlight') | Yes | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Standard | Auto | 512 | 2 | - | -
Extremity Spotlight (extremity) | Yes |  |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Default | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Standard | Auto | 512 | 2 | - | -
SBRT Spine MidDose | Yes |  |  |  |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes (MedDose --> MidDose) |  | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | Default | - | Default | Auto | 512 | 1 | - | -
[5 empty] | Yes (MedDose --> MidDose) | Yes | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
SBRT Spine LowDose | Disabled |  |  |  |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
|  | Disabled |  |  |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Disabled |  |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | Dafault | - | Default | Auto | 512 | 2 | - | -
[5 empty] | Yes | Yes | 125 | 20 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 180 | 0.266 | None | - | Default | Auto | 512 | 2 | - | -
Paediatric Head | Yes |  |  |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Default | Auto | 512 | 1 | - | -
|  | Yes |  |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | None | - | Standard | Auto | 512 | 1 | - | -
|  |  |  | Yes |  | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | Default | - | Default | Auto | 512 | 1 | - | -
[5 empty] | Yes (copy of image gently+modification) | Yes | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | Default | - | Default | Auto
Paediatric Body | Yes |  |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Default | Auto | 512 | 2 | - | -
|  | Yes |  |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
|  |  |  | Yes (copy of image gently) |  | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto
[5 empty] | Yes (copy of image gently) | Yes | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto
KIM LARK Large | Disabled |  |  |  |  | Yes | 140 | 80 | 25 | 15 | 6 | 320mm | 2 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 1000 | 2 | None | - | Default | Smooth | 512 | 2 | - | -
|  | Disabled |  |  |  | Yes | 140 | 80 | 25 | 15 | 6 | 320mm | 2 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 1000 | 2 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Disabled |  |  | Yes | 140 | 80 | 25 | 15 | 6 | 320mm | 2 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 1000 | 2 | None | - | Default | Smooth | 512 | 2 | - | -
No
No
KIM LARK Med | Disabled |  |  |  |  | Yes | 125 | 68 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 850 | 1.394 | None | - | Default | Smooth | 512 | 2 | - | -
|  | Disabled |  |  |  | Yes | 125 | 68 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 850 | 1.394 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Disabled |  |  | Yes | 125 | 68 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 850 | 1.394 | None | - | Default | Smooth | 512 | 2 | - | -
No
No
KIM LARK Small | Disabled |  |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Default | Smooth | 512 | 2 | - | -
|  | Disabled |  |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Disabled |  |  | Yes | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Default | Smooth | 512 | 2 | - | -
No
No
Abdo Spotlight | Yes |  |  |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | - | Requested by RTs. Added to OBK only. Copied Pelvis spotlight mode and then adjusted exposure to match Abdomen, replicating how pelvis spotlight works. JB & EB 2024/02/09
|  | Yes |  |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  | Yes |  |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | -
|  |  |  | Yes |  | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | -
[5 empty] | Yes (copied from 'Pelvis Spotlight') | Yes | 125 | 40 | 20 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | -
Note on O96: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: Not in clinical use
Note on E140: [Threaded comment] Your version of Excel allows you to read this threaded comment; however, any edits to it will get removed if the file is opened in a newer version of Excel. Learn more: https://go.microsoft.com/fwlink/?linkid=870924 Comment: No KIM on BH machines

## Proposed Final
| Current | As of: 24/04/2024
|  |  |  | Acquistion | [15 empty] | Reconstruction
[10 empty] | mGy/100mAs | [5 empty] | deg | s | mAs | cGy
Mode Name | All Linacs | Clinical? | Base mode (copy from to create) | kV | mA/frame | ms/frame | frames/s | deg/s | CTDI Phantom | Dose Factor | Prescan darkfield | Trajectory | Start angle | Fan | No Projections | Proj Increment | Acqusition time | Exposure | CTDI | Default 3D Method | Default 4D Method | Algo | Filter | Matrix size | Slice thick | No. Phases | Phase Bin Width
4D Spotlight | Yes | Yes | Standard | 125 | 40 | 20 | 7 | 3 | 320mm | 1.64 | Yes | Half | Auto | Full | 467 | 0.428 | 67 | 374 | 0.613 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
4D Thorax | Yes | Yes | Standard | 125 | 40 | 20 | 7 | 3 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 840 | 0.429 | 120 | 672 | 0.995 | None | Advanced 4D | 4D | Auto | 384 | 1.5 | 10 | 10
Abdomen | Yes | Yes | Thorax | 125 | 40 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 720 | 1.066 | Standard | - | - | - | 512 | 2 | - | -
Abdo Spotlight | Yes | Yes | Spotlight | 125 | 40 | 20 | 15 | 6 | 320 mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 400 | 0.656 | None | - | Standard | Smooth | 512 | 2 | - | -
Breast 360 | Yes | Yes | Thorax | 125 | 10 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 90 | 0.133 | None | - | Default | Auto | 512 | 2 | - | -
Head | Yes | Yes | Standard | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | Standard | - | Default | Auto | 512 | 2 | - | -
Head SRS | Yes | Yes | Head | 100 | 30 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Full | - | Full (SFOV) | 900 | 0.4 | 60 | 540 | 1.139 | HighRes | - | Default | Auto | 512 | 1 | - | -
Head and Shoulders | Yes | Yes | Thorax | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Default | Auto | 512 | 2 | - | -
Image Gently | Yes | Yes | Standard | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis | Yes | Yes | Standard | 125 | 60 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1080 | 1.598 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Large | Yes | Yes | Pelvis Obese (copy and rename, disable standard clinically | 140 | 75 | 25 | 15 | 6 | 320mm | 2.18 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 1688 | 3.679 | None | - | Standard | Auto | 512 | 2 | - | -
Pelvis Spotlight | Yes | Yes | Spotlight (copy and rename, disable standard clinically) | 125 | 60 | 25 | 15 | 6 | 320mm | 1.64 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 750 | 1.23 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax | Yes | Yes | Standard | 125 | 15 | 20 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 270 | 0.4 | None | - | Standard | Smooth | 512 | 2 | - | -
Thorax Spotlight | Yes | Yes | Spotlight | 125 | 15 | 20 | 15 | 6 | 320mm | 1.32 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.198 | None | - | Default | Auto | 512 | 2 | - | -
Extremity Spotlight | Yes | Yes | Head (direct copy) | 100 | 15 | 20 | 15 | 6 | 160mm | 2.11 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 150 | 0.317 | None | - | Default | Auto | 512 | 2 | - | -
SBRT Spine | Yes | Yes | Thorax | 125 | 40 | 10 | 15 | 6 | 320mm | 1.48 | Yes | Full | - | Half (LFOV) | 900 | 0.4 | 60 | 360 | 0.533 | None | - | Default | Auto | 512 | 2 | - | -
Pediatric Head | Yes | Yes | Image Gently | 80 | 10 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 50 | 0.047 | HiRes | - | Default | Auto | 512 | 1 | - | -
Paediatric Body | Yes | Yes | Image Gently | 80 | 20 | 10 | 15 | 6 | 160mm | 0.94 | Yes | Half | Auto | Full (SFOV) | 500 | 0.4 | 33 | 100 | 0.094 | None | - | Default | Auto | 512 | 2 | - | -
Nonclinical CBCT modes are not on this list (e.g., factory defaults).
Bold indicates change from Varian standard settings

## For printing Apr25
CBCT Modes |  | Current | As of: 24/04/2024
Acquistion
| /mSv | [9 empty] | s | mAs | cGy
Mode Name | Dose | Base mode | Trajectory | kV | mA/f | ms/f | f/s | deg/s | Fan | No Proj | Acq time | Exposure | CTDI
4D Spotlight | 1.6 | Standard | Half | 125 | 40 | 20 | 7 | 3 | Full | 467 | 67 | 374 | 0.613
4D Thorax | 3.3 | Standard | Full | 125 | 40 | 20 | 7 | 3 | Half (LFOV) | 840 | 120 | 672 | 0.995
Abdomen | 2.8 | Thorax | Full | 125 | 40 | 20 | 15 | 6 | Half (LFOV) | 900 | 60 | 720 | 1.066
Abdo Spotlight | 1.2 | Spotlight | Half | 125 | 40 | 20 | 15 | 6 | Full (SFOV) | 500 | 33 | 400 | 0.656
Breast 360 | 0.2 | Thorax | Full | 125 | 10 | 10 | 15 | 6 | Half (LFOV) | 900 | 60 | 90 | 0.133
Head | 0.5 | Standard | Half | 100 | 15 | 20 | 15 | 6 | Full (SFOV) | 500 | 33 | 150 | 0.317
Head SRS | 1.8 | Head | Full | 100 | 30 | 20 | 15 | 6 | Full (SFOV) | 900 | 60 | 540 | 1.139
Head and Shoulders | 0.3 | Thorax | Full | 125 | 15 | 20 | 15 | 6 | Half (LFOV) | 900 | 60 | 270 | 0.4
Image Gently |  | Standard | Half | 80 | 20 | 10 | 15 | 6 | Full (SFOV) | 500 | 33 | 100 | 0.094
Pelvis | 4.2 | Standard | Full | 125 | 60 | 20 | 15 | 6 | Half (LFOV) | 900 | 60 | 1080 | 1.598
Pelvis Large |  | Pelvis Obese | Full | 140 | 75 | 25 | 15 | 6 | Half (LFOV) | 900 | 60 | 1688 | 3.679
Pelvis Spotlight | 2.2 | Spotlight | Half | 125 | 60 | 25 | 15 | 6 | Full (SFOV) | 500 | 33 | 750 | 1.23
Thorax | 1.3 | Standard | Full | 125 | 15 | 20 | 15 | 6 | Half (LFOV) | 900 | 60 | 270 | 0.4
Thorax Spotlight | 0.7 | Spotlight | Half | 125 | 15 | 20 | 15 | 6 | Full (SFOV) | 500 | 33 | 150 | 0.198
Extremity Spotlight | 0.5 | Head | Half | 100 | 15 | 20 | 15 | 6 | Full (SFOV) | 500 | 33 | 150 | 0.317
SBRT Spine | 1.7 | Thorax | Full | 125 | 40 | 10 | 15 | 6 | Half (LFOV) | 900 | 60 | 360 | 0.533
Pediatric Head |  | Image Gently | Half | 80 | 10 | 10 | 15 | 6 | Full (SFOV) | 500 | 33 | 50 | 0.047
Paediatric Body |  | Image Gently | Half | 80 | 20 | 10 | 15 | 6 | Full (SFOV) | 500 | 33 | 100 | 0.094
Nonclinical CBCT modes are not on this list (e.g., factory defaults).
Bold indicates change from Varian standard settings
Effective Dose mSv estimates from PCXMC calculations for generic 80kg/180cm person, scales with body habitus
Documents - ROMP-WSLHD\03 Projects\2023 - CBCT Mode Standardisation

## TB4.1 July25
CBCT Modes TrueBeam 4.1 |  |  |  | Date printed: | 2025-07-30
[10 empty] | s | mAs | mGy
Mode Name | Dose (mSv) | Base mode | Trajectory | kV | mAs Per frame | frame/s | deg/s | Fan | # Proj | Acq time | Exposure | CTDI
4D Spotlight | 1.6 | Standard | Half | 125 | 0.8 | 7 | 3 | Full | 467 | 67 | 373.6 | 6.127
4D Thorax | 3.3 | Standard | Full | 125 | 0.8 | 7 | 3 | Half (LFOV) | 840 | 120 | 671.2 | 9.934
Abdo Spotlight | 1.2 | Spotlight | Half | 125 | 0.8 | 15 | 6 | Full (SFOV) | 500 | 33 | 400.8 | 6.573
Abdomen | 2.8 | Thorax | Full | 125 | 0.8 | 15 | 6 | Half (LFOV) | 900 | 60 | 716 | 10.597
Breast 360 | 0.2 | Thorax | Full | 125 | 0.1 | 15 | 6 | Half (LFOV) | 900 | 60 | 89.5 | 1.325
Extremity Spotlight | 0.5 | Head | Half | 100 | 0.3 | 15 | 6 | Full (SFOV) | 500 | 33 | 150.3 | 3.171
Head | 0.5 | Standard | Half | 100 | 0.3 | 15 | 6 | Full (SFOV) | 500 | 33 | 150.3 | 3.171
Head and Shoulders | 0.3 | Thorax | Full | 125 | 0.3 | 15 | 6 | Half (LFOV) | 900 | 60 | 268.5 | 3.974
Head SRS | 1.8 | Head | Full | 100 | 0.6 | 15 | 6 | Full (SFOV) | 900 | 60 | 537 | 11.331
Image Gently |  | Standard | Half | 80 | 0.2 | 15 | 6 | Full (SFOV) | 500 | 33 | 100.2 | 0.942
Paediatric Body |  | Image Gently | Half | 80 | 0.2 | 15 | 6 | Full (SFOV) | 500 | 33 | 100.2 | 0.942
Pediatric Head |  | Image Gently | Half | 80 | 0.1 | 15 | 6 | Full (SFOV) | 500 | 33 | 50.1 | 0.471
Pelvis | 4.2 | Standard | Full | 125 | 1.2 | 15 | 6 | Half (LFOV) | 900 | 60 | 1074 | 15.895
Pelvis Large |  | Pelvis Obese | Full | 140 | 1.9 | 15 | 6 | Half (LFOV) | 900 | 60 | 1700.5 | 37.071
Pelvis Spotlight | 2.2 | Spotlight | Half | 125 | 1.5 | 15 | 6 | Full (SFOV) | 500 | 33 | 751.5 | 12.325
SBRT Spine | 1.7 | Thorax | Full | 125 | 0.4 | 15 | 6 | Half (LFOV) | 900 | 60 | 358 | 5.298
Thorax | 1.3 | Standard | Full | 125 | 0.3 | 15 | 6 | Half (LFOV) | 900 | 60 | 268.5 | 3.974
Thorax Spotlight | 0.7 | Spotlight | Half | 125 | 0.3 | 15 | 6 | Full (SFOV) | 500 | 33 | 150.3 | 2.465
Bold indicates change from Varian standard settings
Effective Dose mSv estimates from PCXMC calculations for generic 80kg/180cm person, scales with body habitus
TB 4.1 upgrade combines mA & S single value, CTDI now in mGy | [12 empty] | `
Documents - ROMP-WSLHD\03 Projects\2023 - CBCT Mode Standardisation